Perimetric Outcomes of Melbourne Rapid Field Perimetry in Patients with Glaucoma: A Systematic Review and Meta-Analysis

Keywords: perimetry, visual field, telemedicine, glaucoma, tablet, systematic review, meta-analysis

Authors: Andrés Inzunza, MD1,2, Edgar Alejandro Moreno-Diaz, MD3, Deborah Goss, MLS, MA4, David Friedman, MD, MPH, PhD 1,2

Corresponding Author:

David S. Friedman, MD, MPH, PhD

Professor of Ophthalmology, Harvard Medical School

david_friedman@meei.harvard.edu

Financial support: Two authors (AI and EAMD) received a monetary grant from the Mexican National Council of Science and Technology (CONAHCYT/SECIHTI) to support their graduate education. The funding organization had no role in the design or conduct of the research.

Conflict of Interest:

David S. Friedman is a consultant for Théa Pharmaceuticals, AbbVie, Regeneron, and Life Biosciences and has received honoraria for lecturing. David S. Friedman has received research support from Genentech and Perivision. No private or public entity played a role in the design or conduct of the study; collection, management, analysis, or interpretation of the data; preparation, review, or approval of the manuscript; or the decision to submit the manuscript for publication. No other authors have any conflicts of interest to disclose.

Address for reprints:

243 Charles Street

Boston, MA 02114

Abbreviations:

Visual Field = VF

Humphrey Field Analyzer = HFA

Melbourne Rapid Fields = MRF

Swedish Interactive Threshold Algorithm = SITA

Mean deviation = MD

Pattern deviation= PD

Pattern standard deviation = PSD

Ocular hypertension = OHTN

Decibels = dB

1- Harvard Medical School, Boston, MA

2- Department of Ophthalmology, Massachusetts Eye and Ear, Harvard Medical School, Boston, MA

3- Escuela de Medicina, Tecnológico de Monterrey, Guadalajara, México

4- Howe Library, Massachusetts Eye and Ear, Harvard Medical School, Boston, MA

Abstract

Purpose: The Humphrey Field Analyzer (HFA) is the most widely used perimeter in glaucoma, but its cost and low portability limit its use in resource-constrained settings. Melbourne Rapid Fields (MRF) is a low-cost, tablet-based perimeter system that enables at-home testing and testing in low-resource environments. Despite growing evidence that the MRF can obtain reliable estimates of the visual field, no study has synthesized its perimetric performance through meta-analysis and compared it with HFA.

### Methods: Following PROSPERO registration, we conducted a systematic review to compare MRF with HFA in patients with glaucoma. Searches were performed using controlled vocabulary related to perimetry assessment with these devices. Meta-analyses and narrative synthesis were conducted to summarize their performance.

Results: Of 290 studies screened, 9 were included, aggregating data from 777 eyes. Pooled data comparing MRF to HFA, MRF had faster testing (–1.33 minutes), less-negative MD (+2.08 dB), and higher PSD (+0.77 dB). Global agreement with HFA was high (MD ICC = 0.94; MD R² = 0.85; PSD ICC = 0.85). Test–retest reliability was similar between devices (MRF ICC = 0.96; HFA ICC = 0.95). AUCs were comparable for MD (0.84 vs 0.85) but lower for PSD (0.81 vs 0.93). Narrative synthesis indicated greater pointwise variability at intermediate thresholds. MRF had higher and more heterogeneous fixation losses, false positives, and false negatives.

Conclusion: MRF had shorter test durations, high global agreement with HVF, excellent repeatability, and a bias toward less negative global results. MRF showed greater pointwise variability at certain thresholds and worse reliability measures. Despite these differences, MRF returned diagnostic information comparable to HFA.

Précis: MRF had shorter test durations than HFA, high agreement with HFA global indices, modest bias in MD and PSD measurements, and comparable concordance and repeatability. MRF had poorer reliability indices, while maintaining good diagnostic performance.

Introduction

Glaucoma is the leading cause of irreversible vision loss worldwide, affecting an estimated 80 million individuals (1,2). Glaucoma is often asymptomatic in its early stages (3), with studies suggesting that the majority of those affected are unaware of their condition (4). Early detection of glaucoma and glaucoma progression is crucial, as timely intervention can prevent vision loss (5). Nevertheless, the diagnosis of glaucoma remains challenging for healthcare systems (4), as it requires a comprehensive eye exam, intraocular pressure (IOP) measurement, and visual field (VF) assessments to quantify functional vision loss (3).

The Humphrey Field Analyzer (HFA) is the most common method to test the visual field in glaucoma (6), but its use in many settings is limited by the high cost of equipment, the need for trained personnel, and poor portability (7). These barriers are particularly pronounced in low- and middle-income countries (8), where the burden of glaucoma is greatest and access to diagnostic resources is most limited (2). In recent years, tablet-based perimetry has emerged as a new modality for visual field assessment in glaucoma (9). This technology may be particularly valuable for improving access in resource-limited environments, and facilitating home-based testing, offering patients the opportunity for more frequent visual-field assessments (7,10,11).

The Melbourne Rapid Fields (MRF) is a tablet-based perimetry platform that has been validated for assessing visual field function (10). Developed nearly a decade ago (10), MRF has since been tested in patients with glaucoma and other conditions affecting the visual pathways (12,13). Its performance has been studied across multiple geographic regions and in clinical and home settings (7,14). Despite recent growth in the literature on MRF, to our knowledge, no study has synthesized the evidence through meta-analysis and compared its perimetric performance with that of HFA. To address this gap, we conducted a systematic review and meta-analysis of published studies that compared perimetric outcomes and performance between MRF and HFA in patients with glaucoma or suspected of glaucoma.

### Methods

### Following PROSPERO and protocol registration (CRD420251009468), we conducted a systematic review from April 2025 to November 2025 to compare the perimetric performance of MRF with HFA in patients with glaucoma.

### Search Strategy

A systematic review of the literature was designed by the team (AI and DSF) and executed by a medical librarian (DG). The search strategy and its results were documented in accordance with PRISMA guidelines.

Searches were performed on March 11, 2025, across Ovid MEDLINE (1946-), Embase.com (1947-), Web of Science Core Collection (1900-), and Cochrane Library (Wiley 1996-) on March 11, 2025, using controlled vocabulary and keywords focused on the concepts of visual field outcomes using standard automated perimetry versus tablet portable perimetry devices. The full search details are provided in the Appendix. No filters for language, study design, publication date, or country of origin. All references were exported into EndNote 7.8 (Thomson Reuters, 1988-2016) for deduplication, then to Covidence (Veritas Health Innovation) for further deduplication, study screening, selection, and data extraction. The search yielded 527 studies before deduplication, and 289 after deduplication. The study by Chen et al. (2025), published in July 2025, was manually added in November to update the search results, bringing the total to 290 screened studies.

Study inclusion

This review included studies reporting the performance of MRF. Other tablet platforms, such as Online Contrast Circular Perimetry, were excluded due to limited methodological comparability and the use of different normative databases.

We included studies that tested MRF in adults with glaucoma, suspected glaucoma, ocular hypertension, or normal visual fields. We compared the performance of the first-generation MRF with that of the HFA (10). Studies were required to report at least one prespecified outcome of interest and to use the 24-2 SITA Standard algorithm as the comparator. One study included data from SITA Standard and SITA Fast (15); for that study, only the data from SITA Standard were extracted. Publications on patients with glaucoma that compared MRF performance using 30-2 or 10-2 test patterns were excluded to maximize comparability across studies. When a study provided only at-home MRF data, those measurements were included. For studies that reported MRF performance separately for in-clinic and at-home settings, we used the in-clinic data (14).

Studies were imported into COVIDENCE® (Veritas Health Innovation, Melbourne, Australia) for title and abstract screening and full-text review by two independent reviewers (AI and EAMD). Screening of titles and abstracts was performed asynchronously by the two reviewers. Eligible studies underwent a full-text review, with disagreements resolved by written consensus. Data from the included studies were extracted into a centralized external spreadsheet.

We contacted the corresponding authors of 8 of the 9 included studies and submitted 10 data requests, seeking more than 200 additional data points not reported in the published manuscripts. The authors of 6 of the 9 studies provided additional data. Risk of bias was manually assessed by one of the authors (AI) using the QUADAS 2 critical appraisal tools (included in the appendix). Visualization of the risk-of-bias assessment was performed using RStudio (version 2026.01.0+392; Posit Software, Boston, MA, USA) and the “ggplot2” visualization package. No automated tools were used for screening, extraction, or risk of bias appraisal.

Statistical Analysis

We performed meta-analyses to compare the test duration, functional sensitivity, reliability indices, and agreement metrics of MRF and HFA. All analyses were performed using Stata 18 (College Station, TX) (16), with the “meta” and “metan” packages (17). A random-effects model with restricted maximum likelihood (REML) was selected a priori, assuming that the true effect size varies across studies (18). Heterogeneity was quantified using I², τ², H, and Q statistics (19). Standard errors were calculated using standard deviations and the number of eyes as the sample size, then pooled and displayed using forest plots. For outcomes reported as median and IQR, we estimated the mean and standard deviation using the methods described by Wan, Luo, and Tong (20).

The absolute difference in means was used as the effect size to analyze differences in test duration, MD, and PD/PSD between devices. These differences in means were calculated using formulas for paired observations, assuming a high correlation (0.8) (21). This assumption is consistent with the literature (12), and also reflects the expectation that measurements would be strongly correlated since the same participants performed both tests. To further assess this assumption, we conducted sensitivity analyses calculating the difference in means using formulas for unpaired data, and varying the assumed correlation between paired measurements (0.5, 0.8, 0.9)(21).

The correlation coefficients of MD and PSD by MRF and HFA were synthesized using Fisher’s z-transformed correlation coefficients and intraclass correlation coefficients (ICC), with summary effects back-transformed for interpretation (22). The Area Under the Curve (AUC) of Receiver Operating Characteristic (ROC) curves was logit-transformed for meta-analysis. Back-transformed pooled AUCs were used for visual representation in forest plots. Agreement was assessed by meta-analyzing the Bland–Altman MD bias and limits of agreement derived from the study-reported limits of agreement (23). Fixation losses, false positives, and false negatives were meta-analyzed using crude proportions. For six data points that were presented visually but whose numerical values were not reported in the text, we extracted the information using WebPlotDigitizer (24).

To quantitatively explore the possibility of sponsorship bias (25), we performed a meta-regression to examine whether involvement of the technology’s developers (Vingrys or Kong) in data analysis or senior authorship was associated with the reported outcomes (10). In this model, the perimetric outcomes (mean differences in testing time, mean differences in MD and PSD, ICC concordance between MRF and HFA, and test–retest ICC) were used as dependent variables, and the presence or absence of developer involvement was the independent variable.

To ensure transparency and facilitate reproducibility, the original protocol, PROSPERO registration, extracted datasets, risk-of-bias assessments, and complete statistical analysis code have been made publicly available via GitHub (https://github.com/andresinzunza/pop_tag).

Results

A total of 290 studies were screened, 48 underwent full-text review, and 9 were included in the final analysis (Figure 1). Overall, 777 eyes from 638 subjects were included (Table 1). Of these, 644 eyes (83%) had glaucoma, 37 eyes (5%) were classified as glaucoma suspects, and 96 eyes (12%) were normal. The mean age across studies was 61 years (SD: 13.6), and approximately 47% of participants were female. All extracted data were compared between the first version of the MRF technology and the HFA 24-2 using the SITA Standard testing algorithm.

### Testing Time

Seven studies reported testing times, including a total of 692 eyes in this analysis (12,15,26–30). The pooled testing time for HFA was 6.13 minutes per eye (95% CI: 5.89–6.36; I² = 98.9%), compared with 4.82 minutes for MRF (95% CI: 4.3–5.35; I² = 99.8%). The pooled difference in means showed that MRF testing was 1.33 minutes (approximately 1 minute 20 seconds) faster than HFA (95% CI: –1.73 to –0.93; I² = 99.2%) (Supplementary Figure 1).

### MD

Seven studies reported MD, including a total of 559 eyes (12,14,15,27–30). The pooled MD for HFA was –7.15 dB (95% CI: – 9.39 to –4.91; I² = 95.9%), compared with –5.32 dB for MRF (95% CI: –7.17 to –3.47; I² = 94.5%). The pooled mean difference was +2.08 dB (MRF less negative) (95% CI: 1.43–2.73; I² = 61.4%) (Figure 2).

### PSD/PD

Eight studies reported PD/PSD, aggregating information from 692 eyes (12,14,15,26–30). The pooled PSD for HFA was 5.55 dB (95% CI: 4.28–6.28; I² = 96.4%), compared with 6.29 dB for MRF (95% CI: 5.37–7.22; I² = 91%). The pooled mean difference in PSD was +0.77 dB (95% CI: 0.11–1.44; I² = 93.3%) (Figure 3).

### Agreement Analysis (Bland–Altman)

Six studies reported agreement analysis and bias by Bland-Altman analysis, including data from 531 eyes (12,14,15,28–30). The pooled mean bias between MRF and HFA for MD was –2.04 dB ± 2.5dB (HFA – MRF; meaning MRF was less negative). The pooled 95% limits of agreement ranged from –6.96dB to +2.87dB.

### Correlation and Concordance

Six studies reported the correlation and concordance (expressed as R-squared and ICC) of MD obtained by MRF and HFA, aggregating data of 555 eyes (12,14,15,26,28,31). The pooled R² was 0.85 (95% CI: 0.79–0.91; I² = 62%) (Figure 4). Five studies (12,14,15,28,30) reported intraclass correlation coefficients for MD agreement between MRF and HFA, with a total of 428 eyes. The pooled ICC was 0.94 (95% CI: 0.90–0.98; I² = 65%) (Figure 5).

### Zone-Threshold Concordance

### Three studies reported zonal-threshold concordance data, including ICCs for regional mean threshold values in HFA and MRF (12,26,30). Two of the three studies provided results for eight zones (12,26), whereas one study subdivided locations into six zones (30). The pooled data from the two studies that reported the same eight zones (each using its own proprietary testing points, but in the same regions) included 129 eyes. The pooled ICC for Zone 1 (superior–central) was 0.84 (95% CI: 0.76–0.92; I² = 40.2%). For Zone 2 (inferior–central), the pooled ICC was 0.73 (95% CI: 0.64–0.82; I² = 0.0%). The pooled ICC for Zone 3 (superior–peripheral–nasal) was 0.91 (95% CI: 0.88–0.94; I² = 0.0%). For Zone 4 (superior–nasal), the pooled ICC was 0.87 (95% CI: 0.69–1.00; I² = 91.5%). The pooled ICC for Zone 5 (superior–temporal) was 0.80 (95% CI: 0.64–0.97; I² = 81%). For Zone 6 (inferior–temporal), the pooled ICC was 0.79 (95% CI: 0.65–0.94; I² = 73.0%). The pooled ICC for Zone 7 (inferior–nasal) was 0.81 (95% CI: 0.74–0.88; I² = 3.8%). The pooled ICC for Zone 8 (inferior–peripheral–nasal) was 0.91 (95% CI: 0.86–0.95; I² = 28.7%). The pooled ICC for concordance in mean threshold values across the eight regions was 0.85 (95% CI, 0.81–0.88; I² = 78.4%) (Supplementary Figure 2).

### Repeatability

### Five studies reported ICCs of test-retest repeatability of MD measurements of MRF, including a total of 422  eyes (12,26,28,28). Kong et al 2016 (12) assessed immediate test–retest repeatability, Schulz et al. 2018 conducted MRF retesting at 4–6 months, while Prea et al 2018 conducted retesting at 2, 4, and 6 months (15,26). The pooled MRF global test–retest reliability for MD was ICC 0.96 (95% CI: 0.94–0.97; I² = 60.4%). Three studies also evaluated HFA’s test–retest reliability for MD (14,15,28). The pooled HFA test–retest reliability for MD was 0.95 (95% CI: 0.92–0.99; I² = 36%) (Figure 6).

### Zonal Repeatability

### Two studies provided ICCs for regional test-retest values in MRF, pooling data from 175 eyes (12,26). The pooled ICC for Zone 1 (superior–central) was 0.86 (95% CI: 0.82–0.91; I² = 0.0%). For Zone 2 (inferior–central), the pooled ICC was 0.86 (95% CI: 0.81–0.91; I² = 0.0%). The pooled ICC for Zone 3 (superior–peripheral–nasal) was 0.91 (95% CI: 0.82–1.00; I² = 86.3%). For Zone 4 (superior–nasal), the pooled ICC was 0.87 (95% CI: 0.69–1.00; I² = 91.5%). The pooled ICC for Zone 5 (superior–temporal) was 0.82 (95% CI: 0.67–0.98; I² = 86.5%). For Zone 6 (inferior–temporal), the pooled ICC was 0.85 (95% CI: 0.80–0.90; I² = 0.0%). The pooled ICC for Zone 7 (inferior–nasal) was 0.85 (95% CI: 0.90–0.90; I² = 0.0%). The pooled ICC for Zone 8 (inferior–peripheral–nasal) was 0.91 (95% CI: 0.87–0.94; I² = 0.0%). The pooled ICC for regional test-retest values in MRF was 0.87 (95% CI, 0.84–0.90; I² = 74.5%) (Supplementary Figure 3).

### Receiver Operating Characteristic Curves

### Two studies reported AUCs assessing the discriminatory performance of MD and PSD/PD derived from MRF and HFA in glaucomatous and normal eyes, with a total of 188 eyes included in this sub-analysis (26,30). Schulz et al. (2018) contributed 97 eyes (60 glaucomatous, 17 pre-perimetric glaucoma, and 20 age-matched controls), and Chen et al. contributed 91 eyes (58 glaucomatous and 33 normal). MRF had a pooled AUC of 0.84 (0.82–0.85) for MD and 0.81 (0.79–0.82) for PSD, while HFA had a pooled AUC of 0.85 (0.84–0.86) for MD and 0.93 (0.92–0.93) for PSD (Supplementary Figure 4).

### Reliability Indices

Six studies contributed reliability data (12,14,15,26,28,30), including 531 eyes. The pooled proportion of fixation losses was 12% for HFA (95% CI: 4%–21%; I² = 100%) and 26% for MRF (95% CI: 16% – 37%; I² = 99.7%) (Supplementary Figure 5). The pooled false-positive proportion was 2% for HFA (95% CI: 2%–3%; I² = 86.2%) and 7% for MRF (95% CI: 0%–13.4%; I² = 99.5%) (Supplementary Figure 6). The pooled false-negative proportion was 4% for HFA (95% CI: 2%–5%; I² = 94.4%) and 9% for MRF (95% CI: 5%–12%; I² = 88.5%) (Supplementary Figure 7).

Meta-Regression

Meta-regression analyses using developer involvement as the independent variable, and pooled perimetric outcomes as dependent variables found no statistically significant association between the developer involvement on any of the tested outcomes: mean difference in testing time (p = 0.37), mean difference in MD (p = 0.43), mean difference in PSD (p = 0.27), ICC for concordance between MRF and HFA (p = 0.75), and test–retest ICC for MRF (p = 0.73).

### Narrative Synthesis of Point-wise Data

In Kang et al. (2023), a point-by-point analysis was conducted to compare the individual thresholds at 52 locations for HFA and MRF (31). This study showed high agreement across global indices but reported statistically significant differences at 36 of the 52 locations tested. In this study, MRF overestimated light sensitivity in the nasal field and underestimated it in the temporal field, with deviations of up to 6.5 dB in some locations. The authors note that two of the largest discrepancies between HFA and MRF were observed at peripheral points near the physiologic blind spot (31).

Prea et al. (2018) found high concordance and repeatability of global indices between MRF and HFA, but at the pointwise level, repeatability was noisier (15). In this study, pointwise standard deviation increased as thresholds declined, peaking in the 10-20 dB range (with up to a 35% increase in variability) (15). Schulz et al. (2018) similarly found that individual test-retest values were most stable at the extremes (26 to 30 dB and 0 to 6 dB), while intermediate thresholds (12 to 22 dB) showed the greatest dispersion on repeat testing (26).

Prea et al. (2018) showed that MRF produces more unchanged points at retest (15). About 60 percent of MRF locations showed minimal change (pointwise SD 0 to 3.9 dB), compared with only about 10 percent of HFA locations (15). The authors described this as a paradox: despite high global repeatability, MRF exhibits greater pointwise variability at intermediate thresholds during testing. Kong et al. (2016) reported a similar pattern, with ceiling effects at 0 and 30 dB and greater variability at 12 to 17 dB on repeat testing (12). The authors noted that MRF is highly stable for most points, but at mid-range thresholds (10 to 20 dB), when a change occurred, it tended to be larger than with HFA (15). This variability likely arises from the larger steps used by the MRF in this range.

Discussion:

The results of this study are best understood by considering five key architectural differences from HFA. First, this generation of MRF samples more locations (predominantly in the macular region), with 66 test points across 30°, compared with the 52 points of the HFA 24-2 grid (10). Second, the studied MRF version uses a radial test geometry rather than a Cartesian grid (10). Importantly, as a consequence of the radial pattern design, stimulus spacing increases with eccentricity (10). To compensate for this, MRF adds extra testing points in neighboring “gaps” when abnormal responses are detected (12). Third, MRF relies on spatial summation, increasing stimulus size with eccentricity to maintain a constant threshold across the field (creating a “plateau” rather than the classic hill of vision (12), which was intended to stabilize measurements and reduce variability (32). Fourth, MRF employs a seven-level discrete thresholding algorithm, with uneven step sizes (12). Fifth, unlike HFA, the MRF testing procedure requires four fixation changes, and a viewing distance of 33 cm is necessary to avoid corrupting the blind spot (10).

Our findings are consistent with what would be expected from a three-presentation (33), coarser staircase, in which there are faster test times and high global repeatability (particularly at extreme values) with a greater pointwise variability. MRF uses a “minimalist” three-step Bayesian staircase that classifies sensitivity into seven discrete thresholds (0, 6, 12, 17, 22, 26, and 30 dB), derived from the three binary yes/no responses (2³ = eight categories) with the two lowest levels combined due to high variability below 10 dB (10). The algorithm begins with an initial 17dB probe (a value observed by 97.5% of normal patients) (10) and updates the posterior probability distribution after each response. This design speeds testing but reduces granularity; for example, only three levels (12, 17, 22 dB) exist between 12 and 22 dB (10). This “minimalist” design reflects MRF’s original purpose as a portable perimeter intended for frequent home testing, where increased measurement frequency could compensate for reduced pointwise resolution (11).

Given these design features, it is unsurprising that MRF yields different pointwise thresholds, although global indices align closely with HFA values. This outcome also reflects spot scaling: the spot is smaller than Goldman size 3 in the macula, yielding lower thresholds, whereas it is larger in the periphery, yielding higher thresholds. Overall, the size factor is designed to yield HFA-like global values. In our study, MRF demonstrated excellent global concordance, correlation, and AUC, but some performance differences relative to HFA were observed. The higher MD values may be partially explained by the larger stimulus size used to maintain sensitivity across the field. Reliability indices were poorer for MRF; however, these data included at-home measurements, which may have been more susceptible to environmental distractions.

Our study has limitations. Although MRF has been evaluated independently, with findings consistent with those of the technology's developers, a substantial proportion of the included data originated from studies conducted by the developers, raising the possibility of bias owing to conflicts of interest. To address this, we performed a meta-regression to examine whether developer involvement was associated with the reported outcomes. Our analysis found no statistical evidence that such involvement influenced the reported results. Another limitation is the limited availability of meta-analyzable pointwise data, which prevented formal pooling of pointwise outcomes. To mitigate this, we incorporated a narrative synthesis to extract and contextualize clinically relevant, non-meta-analyzable information.

Future research should focus on the impact of the three-presentation coarse staircase on the detection of early glaucomatous defects and quantify the testing frequency at which increased home testing can compensate for a reduced threshold resolution. Another important area for future research is the development of a framework that enables comparison of different portable perimetry technologies while accounting for differences in thresholding strategies, normative databases, and hardware limitations.  Finally, the ability of the MRF to detect change accurately also needs to be confirmed.

In summary, MRF had a small bias in MD and PSD and high agreement with HFA on global indices, with comparable concordance, test–retest reliability, and correlation coefficients. MRF had shorter test durations but exhibited greater pointwise variability, higher rates of fixation losses, false positives, and false negatives. Coupled with its excellent diagnostic performance, the MRF can be a useful perimetric tool in the clinic, and when portable testing is needed.

Table and Figure Legends:

Table 1. Characteristics of included studies and participant populations. *For this study, only SITA Standard data was extracted.

Figure 1. PRISMA flow diagram illustrating the study selection process. A total of 290 records were identified. After removal of duplicates, 288 studies were screened by title and abstract, of which 240 were excluded. Forty-eight full-text articles were assessed for eligibility, and 39 were excluded due to wrong comparator, wrong intervention, or lack of full-text availability. Nine studies met the inclusion criteria and were included in the qualitative and quantitative synthesis.

Figure 2. Forest plot of the pooled difference in means for mean deviation (MD) between Melbourne Rapid Fields (MRF) and Humphrey Field Analyzer (HFA). Individual study estimates and 95% confidence intervals are shown, along with the overall random effects estimate. Positive values indicate that MD values obtained with MRF are less negative (higher dB) compared with HFA.

Figure 3. Forest plot of the pooled difference in means for pattern deviation/pattern standard deviation (PD/PSD) between Melbourne Rapid Fields (MRF) and Humphrey Field Analyzer (HFA). Positive values indicate higher PSD values with MRF compared with HFA.

Figure 4. Forest plot of the pooled r-squared (R²) for mean deviation (MD) between Melbourne Rapid Fields (MRF) and Humphrey Field Analyzer (HFA). Higher R² values indicate stronger correlation and agreement between MRF and HFA measurements. The asymmetric confidence intervals reflect the bounded scale of correlation metrics and the back-transformation of variance-stabilized estimates.

Figure 5: Forest plot of the pooled intraclass correlation coefficient (ICC) for mean deviation (MD) between Melbourne Rapid Fields (MRF) and Humphrey Field Analyzer (HFA). Higher ICC values indicate stronger agreement between MRF and HFA measurements. The asymmetric confidence intervals reflect the bounded scale of correlation metrics and the back-transformation of variance-stabilized estimates.

Figure 6: Forest plot of the pooled intraclass correlation coefficient (ICC) for test–retest repeatability of mean deviation (MD) measurements obtained with Melbourne Rapid Fields (MRF). Higher ICC values indicate greater test–retest reliability of MD measurements.

Supplementary Figure 1. Forest plot comparing per-eye testing time between Melbourne Rapid Fields (MRF) and Humphrey Field Analyzer (HFA). The pooled mean difference indicates that MRF testing is approximately 1.3 minutes (approximately 1 minute and 20 seconds) shorter per eye than HFA. Negative values favor MRF.

Supplementary Figure 2. Forest plot of zonal intraclass correlation coefficients (ICCs) comparing regional mean threshold values between Melbourne Rapid Fields (MRF) and Humphrey Field Analyzer (HFA). Pooled ICCs are shown for each of the eight visual field zones: Zone 1 (superior–central), Zone 2 (inferior–central), Zone 3 (superior–peripheral–nasal), Zone 4 (superior–nasal), Zone 5 (superior–temporal), Zone 6 (inferior–temporal), Zone 7 (inferior–nasal), and Zone 8 (inferior–peripheral–nasal), as well as the overall pooled estimate across zones. Higher ICC values indicate stronger concordance between MRF and HFA regional threshold measurements.

Supplementary Figure 3. Forest plot of zonal intraclass correlation coefficients (ICCs) for test–retest repeatability of regional mean threshold values obtained with Melbourne Rapid Fields (MRF). Pooled ICCs are shown for each of the eight visual field zones: Zone 1 (superior–central), Zone 2 (inferior–central), Zone 3 (superior–peripheral–nasal), Zone 4 (superior–nasal), Zone 5 (superior–temporal), Zone 6 (inferior–temporal), Zone 7 (inferior–nasal), and Zone 8 (inferior–peripheral–nasal), as well as the overall pooled estimate across zones. Higher ICC values indicate greater test–retest repeatability of MRF regional threshold measurements.

Supplementary Figure 4. Forest plot of pooled areas under the receiver operating characteristic curve (AUCs) for the discrimination of glaucomatous versus normal eyes using mean deviation (MD) and pattern standard deviation (PSD) derived from Melbourne Rapid Fields (MRF) and Humphrey Field Analyzer (HFA). Subgroup analyses are shown for MRF MD, HFA MD, MRF PSD, and HFA PSD. Higher AUC values indicate better discriminatory performance. Note: Normal subjects were classified as such by an ophthalmologist following an ophthalmologic examination

Supplementary Figure 5. Forest plot of the pooled proportion (raw proportions) of fixation losses for Humphrey Field Analyzer (HFA) and Melbourne Rapid Fields (MRF). The sub-group comparison assesses differences in fixation loss proportions between devices. Higher proportions indicate a greater frequency of fixation losses.

Supplementary Figure 6. Forest plot of the pooled proportion of false-positive responses for Melbourne Rapid Fields (MRF) and Humphrey Field Analyzer (HFA). Of note, the pooled confidence interval for the MRF subgroup is markedly wide, reflecting substantial between-study heterogeneity and limited precision in this estimation.

Supplementary Figure 7. Forest plot of the pooled proportion of false-negative responses for Melbourne Rapid Fields (MRF) and Humphrey Field Analyzer (HFA). The pooled false-negative proportion is higher for MRF than for HFA.

References:

1.	Sun Y, Chen A, Zou M, Zhang Y, Jin L, Li Y, et al. Time trends, associations and prevalence of blindness and vision loss due to glaucoma: an analysis of observational data from the Global Burden of Disease Study 2017. BMJ Open. 2022 Jan;12(1):e053805.

2.	Tham YC, Li X, Wong TY, Quigley HA, Aung T, Cheng CY. Global Prevalence of Glaucoma and Projections of Glaucoma Burden through 2040. Ophthalmology. 2014 Nov;121(11):2081–90.

3.	Gedde SJ, Vinod K, Wright MM, Muir KW, Lind JT, Chen PP, et al. Primary Open-Angle Glaucoma Preferred Practice Pattern®. Ophthalmology. 2021 Jan;128(1):P71–150.

4.	Soh Z, Yu M, Betzler BK, Majithia S, Thakur S, Tham YC, et al. The Global Extent of Undetected Glaucoma in Adults. Ophthalmology. 2021 Oct;128(10):1393–404.

5.	Heijl A, Leske MC, Bengtsson B, Hyman L, Bengtsson B, Hussein M. Reduction of Intraocular Pressure and Glaucoma Progression. ARCH OPHTHALMOL. 2002;120.

6.	Phu J, Khuu SK, Nivison-Smith L, Kalloniatis M. Standard automated perimetry for glaucoma and diseases of the retina and visual pathways: Current and future perspectives. Prog Retin Eye Res. 2025 Jan;104:101307.

7.	Garba F, Kyari F, Burton MJ, Adler SH, Khan Z, Bello M, et al. Portable devices for the diagnosis of glaucoma: a scoping review. BMJ Open. 2025 Oct;15(10):e105681.

8.	Meethal NSK, Sisodia VPS, George R, Khanna RC. Barriers and Potential Solutions to Glaucoma Screening in the Developing World: A Review. J Glaucoma. 2024 Aug;33(8S):S33.

9.	Garba F, Kyari F, Burton MJ, Adler SH, Khan Z, Bello M, et al. Portable devices for the diagnosis of glaucoma: a scoping review. BMJ Open. 2025 Oct;15(10):e105681.

10.	Vingrys AJ, Healey JK, Liew S, Saharinen V, Tran M, Wu W, et al. Validation of a Tablet as a Tangent Perimeter. Transl Vis Sci Technol. 2016 July 12;5(4):3.

11.	Anderson AJ, Bedggood PA, George Kong YX, Martin KR, Vingrys AJ. Can Home Monitoring Allow Earlier Detection of Rapid Visual Field Progression in Glaucoma? Ophthalmology. 2017 Dec;124(12):1735–42.

12.	Kong YXG, He M, Crowston JG, Vingrys AJ. A Comparison of Perimetric Results from a Tablet Perimeter and Humphrey Field Analyzer in Glaucoma Patients. Transl Vis Sci Technol. 2016 Nov 3;5(6):2.

13.	Prea SM, Kong GYX, Guymer RH, Sivarajah P, Baglin EK, Vingrys AJ. The Short-Term Compliance and Concordance to in Clinic Testing for Tablet-Based Home Monitoring in Age-Related Macular Degeneration. Am J Ophthalmol. 2022 Mar;235:280–90.

14.	Prea SM, Vingrys AJ, Kong GYX. Test Reliability and Compliance to a Twelve-Month Visual Field Telemedicine Study in Glaucoma Patients. J Clin Med. 2022 July 25;11(15):4317.

15.	Prea SM, Kong YXG, Mehta A, He M, Crowston JG, Gupta V, et al. Six-month Longitudinal Comparison of a Portable Tablet Perimeter With the Humphrey Field Analyzer. Am J Ophthalmol. 2018 June;190:9–16.

16.	Stata C. Stata 18. StataCorp LLC.; 2023. (Stata).

17.	StataCorp. Stata 18 Meta-Analysis Reference Manual. College Station, TX: Stata Press; 2023.

18.	Riley RD, Higgins JPT, Deeks JJ. Interpretation of random effects meta-analyses. BMJ. 2011 Feb 10;342:d549.

19.	Higgins JPT, Thompson SG, Deeks JJ, Altman DG. Measuring inconsistency in meta-analyses. BMJ. 2003 Sept 6;327(7414):557–60.

20.	Luo D, Wan X, Liu J, Tong T. Optimally estimating the sample mean from the sample size, median, mid-range, and/or mid-quartile range. Stat Methods Med Res. 2018 June 1;27(6):1785–805.

21.	Chapter 6: Choosing effect measures and computing estimates of effect | Cochrane [Internet]. [cited 2025 Dec 4]. Available from: https://www.cochrane.org/authors/handbooks-and-manuals/handbook/current/chapter-06#section-6-5-2

22.	van Aert RCM. Meta-analyzing partial correlation coefficients using Fisher’s z transformation. Res Synth Methods. 2023 Sept;14(5):768–73.

23.	Tipton E, Shuster J. A Framework for the Meta-Analysis of Bland-Altman Studies Based on a Limits of Agreement Approach. Stat Med. 2017 Oct 15;36(23):3621–35.

24.	Drevon D, Fursa SR, Malcolm AL. Intercoder Reliability and Validity of WebPlotDigitizer in Extracting Graphed Data. Behav Modif. 2017 Mar 1;41(2):323–39.

25.	Jefferson T. Sponsorship bias in clinical trials: growing menace or dawning realisation? J R Soc Med. 2020 Apr;113(4):148–57.

26.	Schulz AM, Graham EC, You Y, Klistorner A, Graham SL. Performance of iPad-based threshold perimetry in glaucoma and controls. Clin Experiment Ophthalmol. 2018;46(4):346–55.

27.	Kumar H, Thulasidas M. Comparison of Perimetric Outcomes from Melbourne Rapid Fields Tablet Perimeter Software and Humphrey Field Analyzer in Glaucoma Patients. J Ophthalmol. 2020;2020(1):8384509.

28.	Prea SM, Kong GYX, Guymer RH, Vingrys AJ. Uptake, Persistence, and Performance of Weekly Home Monitoring of Visual Field in a Large Cohort of Patients With Glaucoma. Am J Ophthalmol. 2021 Mar 1;223:286–95.

29.	Prince J, Thompson A, Mwanza JC, Tolleson-Rinehart S, Budenz DL. Glaucoma Screening Using an iPad-Based Visual Field Test in a West African Population. Ophthalmol Glaucoma. 2022 May;5(3):275–83.

30.	Chen DF, Zhou W, Lu X, Wang T, Wang C, Wang Y, et al. Comparative Analysis of Melbourne Rapid Fields Web-Browser Perimeter and Humphrey Field Analyzer in Chinese Glaucoma Patients. Transl Vis Sci Technol. 2025 July 1;14(7):18.

31.	Kang J, De Arrigunaga S, Freeman SE, Zhao Y, Lin M, Liebman DL, et al. Comparison of Perimetric Outcomes from a Tablet Perimeter, Smart Visual Function Analyzer, and Humphrey Field Analyzer. Ophthalmol Glaucoma. 2023 Sept 1;6(5):509–20.

32.	Bedggood P, Prea SM, Kong YXG, Vingrys AJ. Scaling the size of perimetric stimuli reduces variability and returns constant thresholds across the visual field. J Vis. 2021 Oct 7;21(11):2.

33.	King-Smith PE, Grigsby SS, Vingrys AJ, Benes SC, Supowit A. Efficient and unbiased modifications of the QUEST threshold method: Theory, simulations, experimental evaluation and practical implementation. Vision Res. 1994 Apr 1;34(7):885–912.

Registration: PROSPERO registration CRD420251009468.

Acknowledgments: We are grateful to Prof. Vingrys for sharing his expertise, providing comments, and making available raw data from his group’s studies, which were independently analyzed by our group. We also thank Stefania Papatheodorou and Tobias Elze for their time, expertise, and methodological guidance.

Availability of data and statistical code: The extracted datasets, risk-of-bias assessments, and complete statistical analysis code have been made publicly available (https://github.com/andresinzunza/pop_tag).

Declaration of generative AI and AI-assisted technologies in the writing process: During the preparation of this work, the authors used ChatGPT 4.0-5.2 (OpenAI), Harvard University’s Sanbox (Harvard’s secure Generative AI environment), and Grammarly AI (© 2025 Grammarly, Inc.) to improve the readability and language of the manuscript. ChatGPT 4.0-5.2 (OpenAI) was used for the creation and debugging of the statistical code. After using these tools, the authors reviewed and edited the content as needed and take full responsibility for the content of the article.